LCMS检测条件：

Instrument: SHIMADZU LCMS-2020
Column: C18 4.6*50mm, 3.5um

Mobile phase: A=Water V.03 HCOOH  B=Acetonitrile
Gradient: 25%-25% B in 1min,25%-100% B in 3min，100%-100% B in 0.5min
Flow Rate: 0.8ml/min

Column temperature:35°C

GCMS检测条件 

仪器为Agilent 7890 

检测条件

1、Column: HP-5MS  5% Phenyl Methyl Silox　

   Length: 30M　

   Internal Dia.:0.25mm

   Film thickness: 0.25μm　

   Carrier gas: He

     Flow: 1mL/min

2、Oven Temperature

Oven(initial): 35℃

   Initial hold: 2min

   Rate: 5℃/min

   Oven: 60℃

   Hold: 0min

   Rate: 30℃/min

   Oven(final): 290℃

   Final hold:40min

3、Sample: 

   Injection size: 0.2μL

